Compliance chart Thermo42i NO2 analsyer with ISO 7996-1985 Ambient air -- Determination of the mass concentration of nitrogen oxides - Chemiluminescence method

|  | Relevant section in Standard | Spec. in ISO 7996 | Thermo 42i NOx/NO/NO2 Analyser | Compliant Yes/No |
| --- | --- | --- | --- | --- |
| Methodology | s.3 | Chemiluminescent Method | Chemiluminescent Method | Yes |
| Principal of Operation & reaction | s.3 & s.4 | Principal & reaction laid out in s.3 & s.4 of standard | Principal of Operation and Reaction laid out in Chapter 1 operating manual | Yes |
| Apparatus |  | Attached Schematic | See Attached Schematic | Yes |
| Sampling Line | S.5.1 | Inert to NOx/NO/NO2 | PTFE sample line | Yes |
| Particle Filter | S.5.2 | 5 micrometer pore size | 5 micron zylon PTFE filter used | Yes |
| Control and regulation of sample flow rate | S.5.3 | Sample flow rate to be maintained +/- 2% of manufacturers spec. | Manu. Spec sample flow 500 +/- 50 cc/min. Fault light indicates problems with sample flow | Yes |
| Converter | S.5.4 | A heated furnace maintained at a constant temperature | The No2 to No converter heats molybdenum to approximately 325C. | Yes |
| Ozone Generator | s.5.5 | Ozone generated by a high voltage silent electric discharge. Flow rate kept constant. | Ozone generated by drying ambient air, passing the air between two electrodes that have a large oscillating electric field generated by a high voltage transformer. Solid state flow meter controlling ozone flow. | Yes |
| Reaction Cell | s.5.6 | Constructed of inert material and heated slightly to avoid condensation | Reaction cell located in sensor module. Inbuilt cell heater/thermistor | Yes |
| Optical Filter | S.5.7 | Filter will remove all radiation at wavelengths below 600nm | Optical filter installed on equipment to remove radiation. | Yes |
| Photomultiplier tube | S.5.8 | Housed within a thermostatically controlled refrigerated container | The temperature of the PMT is closely controlled by a dedicated proportional temperature controller. The nominal set-point is -3. Fault light indicates problems with PMT T. | Yes |
| Ozone Filter | S.5.9 | Removal of ozone through activated charcoal | Charcoal ozone scrubber installed | Yes |
| Sampling pump | s.5.10 | Capable of operating under the pressure conditions required for the reaction chamber | Pump provides 1 slpm at 5"Hg-A | Yes |
| Procedure | s.6 | General statement concerning operation of apparatus | Analysers operated as per manufacturer spec. | Yes |
| Determination |  | The analyser will be connected to a suitable recorder or datalogger/computer, taking readings for NO/Nox/NO2 | The analysers are attached to code activated switch or datalogger with NOx/NO/NO2 outputs | Yes |
| NO/NOx cal gas Mixture |  | Periodical check against another independent cal. cylinder | See attached EPA intercalibration result | Yes |

Signed: _______________________________________		Date: ___________